## Slide 1: NeuroDOT
Tutorial:
Image Reconstruction

## Slide 2
NeuroDOT Set-Up - Checklist
Before starting any tutorial, please make sure you’ve done the following
Set your path so that all required folders and subfolders are included
  Folders w/subfolders to include
    NeuroDOT
    Any other toolbox used for whatever task
    Your personal directory
  Folders to include
    Output directory
Set your output directory so that you can save files to the desired location
  Change your current directory to your output directory
    cd(outputdir)

## Slide 3
NeuroDOT Set-Up – Output Directory
Once you’ve set your path, the next thing you should do is set your output directory.
  The output directory must be located within your path
  If it isn’t already in your path, you must add it to your path
You can set your own path to an output directory in the line shown above
An output directory specifies the location where files will be saved
  MATLAB, by default, will save files to the current directory
The following line of code will change your current directory to the output directory that you just specified 
Now you can save outputs

## Slide 4: Diffuse Optical Tomography
In diffuse optical tomography (DOT), arrays of optical sources and detectors are used to perform functional neuroimaging experiments on task-based or resting state modes of human brain function.
Analyses are broken into several pipelines:
  B: modeling of the tissue shape, optical property distribution, and source/detector locations
  C: modeling of the light emission, diffusion, and detection through the head
  D: preprocessing of the raw source-detector measurements
  E: reconstruction and spectroscopy of the preprocessed data and light model into a functional neuroimaging volume
  F: post-processing analysis of these results.

## Slide 5: Image Reconstruction
This tutorial follows the NeuroDOT_ImageReconstruction_Script.m Matlab file which can be found in the Documentation/Scripts folder.
It expands on the full tutorial, NeuroDOT_Full_Processing_Script.m, and assumes the necessary sensitivity matrices are available
The tutorial will:
  Process source-detector measurements
  Perform image reconstruction using a light model sensitivity matrix A
  Visualize the reconstructed data
We will use data from a rotating-wedge retinotopy experiment
Start by loading this data set:
load('NeuroDOT_Data_Sample_CCW1.mat'); % data, info, flags

## Slide 6: Flow Chart for Pre-processing and Image Reconstruction
DOT Acquisition
Log-mean
Detect noisy channels
Band-pass filter
Superficial signal regression
Resample
Low-pass filter
Reconstruction
Spectroscopy
Spatial normalization
Light
Model

## Slide 7: Pre-processing
Pre-processing generally follows this pipeline:
  lmdata = logmean(data); % Logmean Light Levels
  info = FindGoodMeas(lmdata, info, 0.075); % Detect Noisy Channels
  lmdata = detrend_tts(lmdata); % Detrend Data
  lmdata = highpass(lmdata, .02, info.system.framerate); % HPF
  lmdata = lowpass(lmdata, 1, info.system.framerate); % LPF 1
  hem = gethem(lmdata, info); % Superficial Signal Regression
  [lmdata, ~] = regcorr(lmdata, info, hem);
  lmdata = lowpass(lmdata, 0.5, info.system.framerate); % LPF 2
  [lmdata, info] = resample_tts(lmdata, info, 1, 1e-5);
  % Block Average the measurement data
  badata = BlockAverage(lmdata, … info.paradigm.synchpts(info.paradigm.Pulse_2), dt);
  badata=bsxfun(@minus,badata,mean(badata,2));
Plots and images on right show differential measurements, computed as the logarithm of the ratio of each signal and its temporal average
Preprocessing comprises filters, processing, and averaging in order to reduce noise and physiologic components of the signals to highlight response to the stimulus
For a more in-depth exploration of pre-processing parameter settings, please see the Tutorial on Pre-processing.
After
Before
Block-averaged

## Slide 8: Reconstruction & Spectroscopy
DOT Acquisition
Log-mean
Detect noisy channels
Band-pass filter
Superficial signal regression
Resample
Low-pass filter
Reconstruction
Spectroscopy
Light
Model
Absorption image volumes are reconstructed from the processed measurements based on a regularized inversion of the A matrix (for details, see, e.g., Eggebrecht et al., Neuroimage, 2012)
  %% Reconstruct absorption, each wavelength
  if ~exist('A', 'var')
  load('A_AdultV24x28.mat') % Contains A-matrix, info.
  end
  for j = 1:2 % each wavelength
  keep = (info.pairs.WL == j) & (info.pairs.r2d <= 40) & info.MEAS.GI;
  % Invert A-Matrix
  iA = Tikhonov_invert_Amat(A(keep, :), 0.01, 0.1);
  % Smooth Inverted A-Matrix
  iA_smoothed = smooth_Amat(iA, dim, 3);
  % Reconstruct Image Volume
  cortex_mu_a(:, :, j) = reconstruct_img(preprocessed(keep, :), iA_smoothed);
  end
  Image volumes store estimated absorption values as voxels by time by wavelength.
  Hemoglobin images are computed from the wavelength-dependent absorption images
  %% Spectroscopy, convert to Hb
  load('E.mat')
  cortex_Hb = spectroscopy_img(cortex_mu_a, E);
  cortex_HbO = cortex_Hb(:, :, 1);
  cortex_HbR = cortex_Hb(:, :, 2);
  cortex_HbT = cortex_HbO + cortex_HbR;

## Slide 9: Spatial normalization and visualization
DOT Acquisition
Log-mean
Detect noisy channels
Band-pass filter
Superficial signal regression
Resample
Low-pass filter
Reconstruction
Spectroscopy
Spatial normalization
Light
Model
As in the full tutorial, the reconstructed images are visualized on a spatially registered anatomical background
  [MNI,infoB]=LoadVolumetricData('Segmented_MNI152nl_on_MNI111',[],'4dfp');
  MNI_dim = affine3d_img(MNI,infoB,A.info.tissue.dim,eye(4),'nearest');
  HbO_T1 = affine3d_img(cortex_HbO, A.info.tissue.dim, infoB, eye(4));
  Block average and visualize
  badata_HbO = BlockAverage(cortex_HbO,info.paradigm.synchpts(info.paradigm.Pulse_2), dt);
  badata_HbO=bsxfun(@minus,badata_HbO,badata_HbO(:,1));
  badata_HbOvol = Good_Vox2vol(badata_HbO,A.info.tissue.dim);
  Params.Scale=0.8*max(abs(badata_HbOvol(:)));
  Params.Th.P=0; %positive threshold
  Params.Th.N=-Params.Th.P; %negative threshold
  PlotSlicesTimeTrace(MNI_dim,A.info.tissue.dim,Params,badata_HbOvol,info)
  The results below show the reconstructed oxyhemoglobin (HbO) volume.

## Slide 10: Visualization of Absorption Volumes
Reconstructed absorption volumes at 750 and 850 nm show sensitivity of the 850 nm wavelength to HbO and 750 nm wavelength to HbR.
  Images are shown with a threshold of 20% of the maximum amplitude:
  Params.Scale=0.8*max(abs(badata_HbOvol(:)));
  Params.Th.P=0.2*Params.Scale;
  Params.Th.N=-Params.Th.P;
  Params.Cmap='jet';
  PlotSlicesTimeTrace(MNI_dim,A.info.tissue.dim,Params,badata_HbOvol,info)
HbO
- 750 nm
- 850 nm

## Slide 11: Hemoglobin volumes - HbO, HbR, HbT
  Oxy-, deoxy-, and total hemoglobin estimates can also be visualized
HbO
HbR
HbT

## Slide 12: Spatial Smoothing
Spatial smoothing is implemented with a 3-dimensional Gaussian filter. The third parameter dictates the degree of smoothing via the filter width, in voxels
  % Smooth Inverted A-Matrix
  iA_smoothed = smooth_Amat(iA, dim, 3);
  Increasing smoothing generally gives improved signal to noise ratio at the expense of resolution
HbO – Smoothed with 5 voxel filter
HbO – Smoothed with 3 voxel filter
HbO – Not smoothed

## Slide 13: Reconstruction Parameters
Reconstruction via regularized inversion is controlled by two parameters
  iA = Tikhonov_invert_Amat(A.A(keep, :), 0.01, 0.1); % Invert A-Matrix
  The first parameter controls a tradeoff between image resolution and noise, similar to spatial smoothing. The typical value for this parameter is
  The second parameter controls spatially variant regularization. A lower value increases reconstruction amplitude in low-sensitivity voxels, e.g., deeper voxels. Typical value is
HbO – ,
HbO – ,
HbO – , 05

## Slide 14
That’s It (For Now)
Congratulations! You have finished the NeuroDOT Image Reconstruction Tutorial
You are now ready to start the other NeuroDOT Tutorials
Recommended order for the tutorials is:
  Generating a Light Model
  Preprocessing
  Image Reconstruction
  Full Data Processing
  Generating a Light Model PAD Adult 96x92
If you have any questions about the material presented in this tutorial, please consult the NeuroDOT User Manual on the WUSTL Optical Radiology Lab (ORL) GitHub page https://github.com/WUSTL-ORL/NeuroDOT_Beta in the documentation folder, or contact either: Adam Eggebrecht: aeggebre@wustl.edu, Emma Speh: espeh@wustl.edu, Ari Segel: ari@wustl.edu, or Jason Trobaugh: jasont@wustl.edu